CEMRR Gait Pipeline

Hand Calculations — 10 Gait Cycles

CP Hemiplegia (Left affected) — Raw Sensor Data to GRI

Center of Excellence in Medical Robotics and Rehabilitation Research (CEMRR)

Nazarbayev University, Astana, Kazakhstan

Patient and System Parameters

| Parameter | Value |
| --- | --- |
| Diagnosis | CP Hemiplegia (Left side affected) |
| Body weight | 32 kg |
| Leg length (L) | 0.62 m |
| Treadmill belt speed | 0.50 m/s |
| Gait cycles measured | 10 |
| Sensors | 4 encoders + 4 Futek® load cells + 2×16 insoles |
| Load cell moment arm | 0.25 m |
| Insole sample rate | 200 Hz |

1. Raw Sensor Data — 10 Gait Cycles

1.1 Encoder Data — Joint Angles (°)

Each value is the peak flexion angle measured by the encoder at that motor during that gait cycle (C1-C10). Hip angle is used to compute step length; knee angle is used to assess joint ROM.

| Parameter | C1 | C2 | C3 | C4 | C5 | C6 | C7 | C8 | C9 | C10 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| θ_hip L (°) | 24.2 | 23.8 | 24.5 | 24.1 | 23.9 | 24.3 | 24.6 | 24.0 | 23.7 | 24.4 |
| θ_hip R (°) | 36.1 | 36.4 | 35.9 | 36.3 | 36.0 | 36.5 | 36.2 | 36.1 | 36.3 | 36.0 |
| θ_knee L (°) | 51.2 | 50.8 | 51.5 | 51.0 | 50.9 | 51.3 | 51.6 | 51.1 | 50.7 | 51.4 |
| θ_knee R (°) | 58.4 | 58.7 | 58.2 | 58.6 | 58.3 | 58.8 | 58.5 | 58.4 | 58.6 | 58.3 |

1.2 Insole Gait Events — Heel Strike (HS) and Toe-Off (TO) Times (s)

HS detected when heel sensors (13–16) sum crosses 5% of max pressure (upward crossing). TO detected when total insole sum drops below 5% threshold. 11 HS events give 10 complete strides.

| Event | E1 | E2 | E3 | E4 | E5 | E6 | E7 | E8 | E9 | E10 | E11 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| HS_L (s) | 0.000 | 1.284 | 2.571 | 3.856 | 5.143 | 6.428 | 7.715 | 9.000 | 10.289 | 11.574 | 12.861 |
| HS_R (s) | 0.642 | 1.926 | 3.211 | 4.498 | 5.783 | 7.068 | 8.355 | 9.640 | 10.927 | 12.212 | 13.499 |

| Event | T1 | T2 | T3 | T4 | T5 | T6 | T7 | T8 | T9 | T10 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| TO_L (s) | 0.800 | 2.084 | 3.371 | 4.656 | 5.943 | 7.228 | 8.515 | 9.800 | 11.089 | 12.374 |
| TO_R (s) | 1.380 | 2.664 | 3.949 | 5.236 | 6.521 | 7.806 | 9.093 | 10.378 | 11.665 | 12.950 |

1.3 Futek® Load Cell Data — Peak Force per Cycle (N)

Each load cell is placed in series with the exoskeleton actuator. The passive reference is recorded at calibration with the patient fully relaxed. The difference (F_passive − F_session) × moment_arm = τ_human, the torque the patient contributes.

| Sensor | C1 | C2 | C3 | C4 | C5 | C6 | C7 | C8 | C9 | C10 | F_passive |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| F_hip_L (N) | 3.2 | 3.5 | 3.1 | 3.4 | 3.3 | 3.6 | 3.2 | 3.5 | 3.3 | 3.4 | 5.1 |
| F_hip_R (N) | 2.8 | 3.0 | 2.9 | 2.8 | 3.1 | 2.9 | 3.0 | 2.8 | 3.0 | 2.9 | 4.6 |
| F_knee_L (N) | 2.4 | 2.6 | 2.3 | 2.5 | 2.4 | 2.7 | 2.4 | 2.5 | 2.3 | 2.6 | 4.0 |
| F_knee_R (N) | 2.0 | 2.1 | 2.0 | 2.2 | 2.1 | 2.0 | 2.2 | 2.1 | 2.0 | 2.1 | 3.7 |

1.4 Insole Mean Pressure — L1–L16 and R1–R16 (% Body Weight)

| Foot | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Left | 9.1 | 7.8 | 6.8 | 8.5 | 11.2 | 10.8 | 11.9 | 11.4 | 11.8 | 7.2 | 5.9 | 7.1 | 9.5 | 14.2 | 12.9 | 10.8 |
| Right | 9.8 | 5.5 | 3.0 | 15.8 | 12.4 | 13.9 | 9.8 | 7.0 | 7.5 | 5.9 | 5.8 | 6.8 | 7.0 | 5.4 | 8.3 | 5.4 |

2. Hand Calculations — Step by Step

2.1 Stride Time, Stance Time, Swing Time, Double Support

Formulas:

t_stride_L(i) = HS_L(i+1) − HS_L(i)    [full gait cycle time]

t_stance_L(i) = TO_L(i)   − HS_L(i)    [foot ON ground]

t_swing_L(i)  = HS_L(i+1) − TO_L(i)    [foot IN air]

t_ds(i)       = min(TO_L(i),TO_R(i)) − max(HS_L(i),HS_R(i))  [both feet on ground]

Worked example — Cycle 1:

t_stride_L(1) = HS_L(2) − HS_L(1) = 1.284 − 0.000 = 1.284 s

t_stance_L(1) = TO_L(1) − HS_L(1) = 0.800 − 0.000 = 0.800 s

t_swing_L(1)  = HS_L(2) − TO_L(1) = 1.284 − 0.800 = 0.484 s

Check: 0.800 + 0.484 = 1.284 s = t_stride_L(1) ✓

t_stride_R(1) = HS_R(2) − HS_R(1) = 1.926 − 0.642 = 1.284 s

t_stance_R(1) = TO_R(1) − HS_R(1) = 1.380 − 0.642 = 0.738 s

t_swing_R(1)  = HS_R(2) − TO_R(1) = 1.926 − 1.380 = 0.546 s

t_ds(1) = min(TO_L(1),TO_R(1)) − max(HS_L(1),HS_R(1))

= min(0.800, 1.380) − max(0.000, 0.642)

= 0.800 − 0.642 = 0.158 s

All 10 cycles — per-cycle table:

| C | t_str_L | t_st_L | t_sw_L | t_str_R | t_st_R | t_sw_R | t_ds | %str_L | %str_R | DSR_i |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | 1.2840 | 0.8000 | 0.4840 | 1.2840 | 0.7380 | 0.5460 | 0.1580 | 62.3% | 57.5% | 12.31% |
| 2 | 1.2870 | 0.8000 | 0.4870 | 1.2850 | 0.7380 | 0.5470 | 0.1580 | 62.2% | 57.4% | 12.28% |
| 3 | 1.2850 | 0.8000 | 0.4850 | 1.2870 | 0.7380 | 0.5490 | 0.1600 | 62.3% | 57.3% | 12.45% |
| 4 | 1.2870 | 0.8000 | 0.4870 | 1.2850 | 0.7380 | 0.5470 | 0.1580 | 62.2% | 57.4% | 12.28% |
| 5 | 1.2850 | 0.8000 | 0.4850 | 1.2850 | 0.7380 | 0.5470 | 0.1600 | 62.3% | 57.4% | 12.45% |
| 6 | 1.2870 | 0.8000 | 0.4870 | 1.2870 | 0.7380 | 0.5490 | 0.1600 | 62.2% | 57.3% | 12.43% |
| 7 | 1.2850 | 0.8000 | 0.4850 | 1.2850 | 0.7380 | 0.5470 | 0.1600 | 62.3% | 57.4% | 12.45% |
| 8 | 1.2890 | 0.8000 | 0.4890 | 1.2870 | 0.7380 | 0.5490 | 0.1600 | 62.1% | 57.3% | 12.41% |
| 9 | 1.2850 | 0.8000 | 0.4850 | 1.2850 | 0.7380 | 0.5470 | 0.1620 | 62.3% | 57.4% | 12.61% |
| 10 | 1.2870 | 0.8000 | 0.4870 | 1.2870 | 0.7380 | 0.5490 | 0.1620 | 62.2% | 57.3% | 12.59% |
| MEAN | 1.2861 | 0.8000 | 0.4861 | 1.2857 | 0.7380 | 0.5477 | 0.1598 | 62.2% | 57.4% | 12.43% |

2.2 Cadence

Formula:

Cadence = 60 / mean(t_step)   where t_step = time between any two consecutive HS events

All HS events combined (both feet, sorted):

HS events: 0.000s, 0.642s, 1.284s, 1.926s, 2.571s, 3.211s, 3.856s, 4.498s, 5.143s, 5.783s, 6.428s, 7.068s, 7.715s, 8.355s, 9.000s, 9.640s, 10.289s, 10.927s, 11.574s, 12.212s, 12.861s, 13.499s

Step times: 0.642s, 0.642s, 0.642s, 0.645s, 0.640s, 0.645s, 0.642s, 0.645s, 0.640s, 0.645s, 0.640s, 0.647s, 0.640s, 0.645s, 0.640s, 0.649s, 0.638s, 0.647s, 0.638s, 0.649s, 0.638s

Mean step time = (0.642 + 0.642 + 0.642 + 0.645 + 0.640 + 0.645 + 0.642 + 0.645 + 0.640 + 0.645 + 0.640 + 0.647 + 0.640 + 0.645 + 0.640 + 0.649 + 0.638 + 0.647 + 0.638 + 0.649 + 0.638)/21

= 13.4990 / 21 = 0.6428 s

Cadence = 60 / 0.6428 = 93.34 steps/min

Normal reference: 110 steps/min  →  patient at 84.9% of normal

2.3 Step Length and Stride Length

Formulas:

L_kin  = 2 × L_leg × sin(θ_hip_max / 2)   [kinematic from encoder]

L_belt = v_belt × t_step                   [belt correction for treadmill]

L_step = L_kin + L_belt                    [actual step length]

L_stride = L_step_L + L_step_R             [full gait cycle distance]

Worked example — Cycle 1, Left foot:

θ_hip_L(C1) = 24.2°

L_kin_L(C1) = 2 × 0.62 × sin(24.2/2°) = 2 × 0.62 × sin(12.100°)

= 2 × 0.62 × 0.2096 = 0.2599 m

t_step_L(C1)= HS_R(1) − HS_L(1) = 0.642 − 0.000 = 0.6420 s

L_belt_L(C1)= v_belt × t_step_L = 0.500 × 0.6420 = 0.3210 m

L_step_L(C1)= 0.2599 + 0.3210 = 0.5809 m

θ_hip_R(C1) = 36.1°

L_kin_R(C1) = 2 × 0.62 × sin(36.1/2°) = 0.3842 m

t_step_R(C1)= HS_L(2) − HS_R(1) = 1.284 − 0.642 = 0.6420 s

L_belt_R(C1)= 0.500 × 0.6420 = 0.3210 m

L_step_R(C1)= 0.3842 + 0.3210 = 0.7052 m

L_stride(C1)= L_step_L + L_step_R = 0.5809 + 0.7052 = 1.2861 m

All 10 cycles — step and stride lengths:

| C | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | MEAN |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| L_kin_L (m) | 0.2599 | 0.2557 | 0.2631 | 0.2589 | 0.2568 | 0.2610 | 0.2642 | 0.2578 | 0.2546 | 0.2620 | 0.2594 |
| L_kin_R (m) | 0.3842 | 0.3873 | 0.3822 | 0.3863 | 0.3832 | 0.3883 | 0.3852 | 0.3842 | 0.3863 | 0.3832 | 0.3850 |
| L_belt_L (m) | 0.3210 | 0.3210 | 0.3200 | 0.3210 | 0.3200 | 0.3200 | 0.3200 | 0.3200 | 0.3190 | 0.3190 | 0.3201 |
| L_belt_R (m) | 0.3210 | 0.3225 | 0.3225 | 0.3225 | 0.3225 | 0.3235 | 0.3225 | 0.3245 | 0.3235 | 0.3245 | 0.3229 |
| L_step_L (m) | 0.5809 | 0.5767 | 0.5831 | 0.5799 | 0.5768 | 0.5810 | 0.5842 | 0.5778 | 0.5736 | 0.5810 | 0.5795 |
| L_step_R (m) | 0.7052 | 0.7098 | 0.7047 | 0.7088 | 0.7057 | 0.7118 | 0.7077 | 0.7087 | 0.7098 | 0.7077 | 0.7080 |
| L_stride (m) | 1.2861 | 1.2865 | 1.2878 | 1.2886 | 1.2824 | 1.2928 | 1.2919 | 1.2865 | 1.2834 | 1.2887 | 1.2875 |

2.4 Stride Time Symmetry and Stride Length Symmetry

Formulas:

SI_stride_time (%) = |μ_stride_L − μ_stride_R| / ((μ_stride_L + μ_stride_R)/2) × 100

SI_step_length (%) = |μ_step_L   − μ_step_R|   / ((μ_step_L   + μ_step_R)/2)   × 100

Stride time symmetry:

μ_stride_L = 1.2861 s

μ_stride_R = 1.2857 s

|1.2861 − 1.2857| = 0.0004 s

mean = (1.2861 + 1.2857)/2 = 1.2859 s

SI_stride_time = 0.0004 / 1.2859 × 100 = 0.0311%

Step length symmetry (SI):

μ_step_L = 0.5795 m

μ_step_R = 0.7080 m

|0.5795 − 0.7080| = 0.1285 m

mean = (0.5795 + 0.7080)/2 = 0.6437 m

SI_step_length = 0.1285 / 0.6437 × 100 = 19.9592%

2.5 Double Support Ratio (DSR)

Formula:

DSR (%) = μ_t_ds / μ_t_stride_L × 100

Worked example — Cycle 1:

t_ds(1) = min(TO_L(1),TO_R(1)) − max(HS_L(1),HS_R(1))

= min(0.800,1.380) − max(0.000,0.642) = 0.800 − 0.642 = 0.158 s

Mean across 10 cycles:

μ_t_ds = (0.1580 + 0.1580 + 0.1600 + 0.1580 + 0.1600 + 0.1600 + 0.1600 + 0.1600 + 0.1620 + 0.1620) / 10

= 1.5980 / 10 = 0.1598 s

DSR = 0.1598 / 1.2861 × 100 = 12.4252%

Normal reference DSR ≈ 20%  →  patient at 12.43%

2.6 Stride Time Coefficient of Variation (CV)

Formula:

CV (%) = σ_stride / μ_stride × 100

σ_stride = √[ Σ(t_stride(i) − μ)² / (n−1) ]

Left limb (affected):

μ_L = (1.2840 + 1.2870 + 1.2850 + 1.2870 + 1.2850 + 1.2870 + 1.2850 + 1.2890 + 1.2850 + 1.2870) / 10

= 12.8610 / 10 = 1.2861 s

Deviations (t_i − μ)²: 0.00000441, 0.00000081, 0.00000121, 0.00000081, 0.00000121, 0.00000081, 0.00000121, 0.00000841, 0.00000121, 0.00000081

Sum = 0.00002090

σ_L = √(0.00002090/9) = √(0.00000232) = 0.001524 s

CV_L = 0.001524 / 1.2861 × 100 = 0.1185%

Right limb:

CV_R = 0.001160 / 1.2857 × 100 = 0.0902%

Normal CV < 3%  →  L = 0.1185% (normal) · R = 0.0902% (normal)

2.7 Joint Torque from Futek® Load Cells

Formulas:

τ_session(i) = F_session(i) × moment_arm     [torque from session load cell]

τ_passive    = F_passive × moment_arm         [calibration: patient fully relaxed]

τ_human      = τ_passive − τ_session          [patient voluntary torque contribution]

= (F_passive − F_session) × moment_arm

PTF          = (F_passive − F_session_mean) / F_passive  [Patient Torque Fraction]

τ_human > 0  → patient ASSISTING motor (desired)

τ_human < 0  → patient RESISTING motor (spasticity)

Note: Unlike motor current estimation, the Futek® load cell directly measures the force at the joint, making τ_human a direct mechanical measurement with no need for a torque constant k_t.

Joint: L hip

F_passive_L_hip = 5.1 N

F_session per cycle: [3.2, 3.5, 3.1, 3.4, 3.3, 3.6, 3.2, 3.5, 3.3, 3.4] N

Mean F_session = (3.2+3.5+3.1+3.4+3.3+3.6+3.2+3.5+3.3+3.4) / 10 = 3.350 N

τ_human = (F_passive − F_session_mean) × moment_arm

= (5.1 − 3.350) × 0.25 m

= 0.4375 Nm

PTF     = (5.1 − 3.350) / 5.1

= 0.3431  (34.31%)

|  | C1 | C2 | C3 | C4 | C5 | C6 | C7 | C8 | C9 | C10 | MEAN |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| F_session (N) | 3.2 | 3.5 | 3.1 | 3.4 | 3.3 | 3.6 | 3.2 | 3.5 | 3.3 | 3.4 | 3.350 |
| τ_human (Nm) | 0.4750 | 0.4000 | 0.5000 | 0.4250 | 0.4500 | 0.3750 | 0.4750 | 0.4000 | 0.4500 | 0.4250 | 0.4375 |

Joint: R hip

F_passive_R_hip = 4.6 N

F_session per cycle: [2.8, 3.0, 2.9, 2.8, 3.1, 2.9, 3.0, 2.8, 3.0, 2.9] N

Mean F_session = (2.8+3.0+2.9+2.8+3.1+2.9+3.0+2.8+3.0+2.9) / 10 = 2.920 N

τ_human = (F_passive − F_session_mean) × moment_arm

= (4.6 − 2.920) × 0.25 m

= 0.4200 Nm

PTF     = (4.6 − 2.920) / 4.6

= 0.3652  (36.52%)

|  | C1 | C2 | C3 | C4 | C5 | C6 | C7 | C8 | C9 | C10 | MEAN |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| F_session (N) | 2.8 | 3.0 | 2.9 | 2.8 | 3.1 | 2.9 | 3.0 | 2.8 | 3.0 | 2.9 | 2.920 |
| τ_human (Nm) | 0.4500 | 0.4000 | 0.4250 | 0.4500 | 0.3750 | 0.4250 | 0.4000 | 0.4500 | 0.4000 | 0.4250 | 0.4200 |

Joint: L knee

F_passive_L_knee = 4.0 N

F_session per cycle: [2.4, 2.6, 2.3, 2.5, 2.4, 2.7, 2.4, 2.5, 2.3, 2.6] N

Mean F_session = (2.4+2.6+2.3+2.5+2.4+2.7+2.4+2.5+2.3+2.6) / 10 = 2.470 N

τ_human = (F_passive − F_session_mean) × moment_arm

= (4.0 − 2.470) × 0.25 m

= 0.3825 Nm

PTF     = (4.0 − 2.470) / 4.0

= 0.3825  (38.25%)

|  | C1 | C2 | C3 | C4 | C5 | C6 | C7 | C8 | C9 | C10 | MEAN |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| F_session (N) | 2.4 | 2.6 | 2.3 | 2.5 | 2.4 | 2.7 | 2.4 | 2.5 | 2.3 | 2.6 | 2.470 |
| τ_human (Nm) | 0.4000 | 0.3500 | 0.4250 | 0.3750 | 0.4000 | 0.3250 | 0.4000 | 0.3750 | 0.4250 | 0.3500 | 0.3825 |

Joint: R knee

F_passive_R_knee = 3.7 N

F_session per cycle: [2.0, 2.1, 2.0, 2.2, 2.1, 2.0, 2.2, 2.1, 2.0, 2.1] N

Mean F_session = (2.0+2.1+2.0+2.2+2.1+2.0+2.2+2.1+2.0+2.1) / 10 = 2.080 N

τ_human = (F_passive − F_session_mean) × moment_arm

= (3.7 − 2.080) × 0.25 m

= 0.4050 Nm

PTF     = (3.7 − 2.080) / 3.7

= 0.4378  (43.78%)

|  | C1 | C2 | C3 | C4 | C5 | C6 | C7 | C8 | C9 | C10 | MEAN |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| F_session (N) | 2.0 | 2.1 | 2.0 | 2.2 | 2.1 | 2.0 | 2.2 | 2.1 | 2.0 | 2.1 | 2.080 |
| τ_human (Nm) | 0.4250 | 0.4000 | 0.4250 | 0.3750 | 0.4000 | 0.4250 | 0.3750 | 0.4000 | 0.4250 | 0.4000 | 0.4050 |

STR — Overall Joint Strength Score:

STR = (PTF_L_hip + PTF_R_hip + PTF_L_knee + PTF_R_knee) / 4

= (0.3431 + 0.3652 + 0.3825 + 0.4378) / 4

= 1.5287 / 4

= 0.3822  (38.22%)

3. Gait Aspect Scores — Normalised [0→1]

Each raw parameter is normalised relative to a baseline (Session 1 measurement) and a normal reference. Score 0 = same as baseline; score 1 = at normal reference level.

| Aspect | Step-by-step calculation | Score | [0→1] |
| --- | --- | --- | --- |
| S — Symmetry (Avg Step Length SI) | Step 1 — Compute SI from mean step lengths: SI = ¦μ_step_L − μ_step_R¦ / ((μ_step_L + μ_step_R)/2) × 100 = ¦0.5795 − 0.7080¦ / ((0.5795 + 0.7080)/2) × 100 = 0.1285 / 0.6437 × 100 = 19.9592% Step 2 — Normalise (SI_baseline = 20.54%): S = 1 − (SI / SI_baseline) = 1 − (19.9592 / 20.54) = 1 − 0.9717 = 0.0283 | 0.0283 | Needs work |
| V — Stability (Stride Time CV) | Step 1 — Compute CV for both limbs from stride times: CV_L = σ_L / μ_stride_L × 100 = 0.001524 / 1.2861 × 100 = 0.1185% CV_R = σ_R / μ_stride_R × 100 = 0.001160 / 1.2857 × 100 = 0.0902% Step 2 — Mean CV: V_raw = (CV_L + CV_R) / 2 = (0.1185 + 0.0902) / 2 = 0.1043% Step 3 — Normalise (CV_baseline = 4.2%): V = 1 − (V_raw / CV_baseline) = 1 − (0.1043 / 4.2) = 1 − 0.0248 = 0.9752 | 0.9752 | Good |
| B — Support / Balance (Double Support Ratio) | Step 1 — Mean double support time and DSR: μ_t_ds = 0.1598 s (mean over 10 cycles) DSR = 0.1598 / 1.2861 × 100 = 12.4252% Step 2 — Normalise (DSR_normal = 20%, DSR_baseline = 36%): B = 1 − (DSR − 20) / (36 − 20) = 1 − (12.4252 − 20) / 16 = 1 − -0.4734 = 1.0000 → capped at 1.000 (DSR already below 20%) | 1.0000 | Good |
| E — Efficiency / Propulsion (Speed, Cadence, Stride Length) | Step 1 — Normalise each sub-parameter: E_speed = v_belt / v_normal = 0.500 / 0.900 = 0.5556 E_cad = cadence / cad_normal = 93.34 / 110 = 0.8485 E_sl = L_stride / SL_normal = 1.2875 / 1.400 = 0.9196 Step 2 — Equal-weight mean: E = (E_speed + E_cad + E_sl) / 3 = (0.5556 + 0.8485 + 0.9196) / 3 = 2.3237 / 3 = 0.7746 | 0.7746 | Good |
| STR — Joint Strength (Patient Torque Fraction) | Step 1 — PTF per joint from Futek® load cells: PTF_L_hip = (5.1 − 3.350) / 5.1 = 0.3431 (34.31%) PTF_R_hip = (4.6 − 2.920) / 4.6 = 0.3652 (36.52%) PTF_L_knee = (4.0 − 2.470) / 4.0 = 0.3825 (38.25%) PTF_R_knee = (3.7 − 2.080) / 3.7 = 0.4378 (43.78%) Step 2 — Mean PTF = STR: STR = (0.3431 + 0.3652 + 0.3825 + 0.4378) / 4 = 1.5287 / 4 = 0.3822 | 0.3822 | Improving |

4. Gait Recovery Index (GRI)

GRI = 0.25·S + 0.15·V + 0.20·B + 0.20·E + 0.20·STR

| w | Aspect | Weight | Rationale |
| --- | --- | --- | --- |
| w_S | S Symmetry | 0.25 | Primary CP hemiplegia impairment — main therapeutic target |
| w_V | V Stability | 0.15 | Treadmill enforces rhythm; lower weight than spatial indices |
| w_B | B Support/Balance | 0.20 | Loading equality — directly targeted by BWS + insole biofeedback |
| w_E | E Efficiency | 0.20 | Functional walking output; clinically meaningful to patient and carers |
| w_STR | STR Joint Strength | 0.20 | Neurological recovery; directly measured by Futek® load cells |
| Σw | — | 1.00 | Weights sum to exactly 1.0 |

Full numerical expansion:

GRI = 0.25 × S    +  0.15 × V    +  0.20 × B    +  0.20 × E    +  0.20 × STR

= 0.25 × 0.0283  +  0.15 × 0.9752  +  0.20 × 1.0000  +  0.20 × 0.7746  +  0.20 × 0.3822

= 0.0071       +  0.1463       +  0.2000       +  0.1549       +  0.0764

| Component | Value |
| --- | --- |
| S (Symmetry) × 0.25 | 0.0071 ← step length SI = 19.96%, SI_baseline = 20.54% |
| V (Stability) × 0.15 | 0.1463 ← CV_L = 0.1185%, CV_R = 0.0902%, baseline = 4.2% |
| B (Support) × 0.20 | 0.2000 ← DSR = 12.43% (already below 20% normal → capped at 1.00) |
| E (Efficiency) × 0.20 | 0.1549 ← speed=0.5556, cad=0.8485, SL=0.9196 |
| STR (Jt Strength) × 0.20 | 0.0764 ← PTF = 0.3822 (Futek® load cells) |
| GRI (total) | GRI = 0.5847 → 58.5% of the way to normal gait |

4.1 Summary of All Calculated Parameters

| Parameter | Left (affected) | Right | Ref |
| --- | --- | --- | --- |
| Stride time (s) | 1.2861 | 1.2857 | ~1.15 |
| Stance time (s) | 0.8000 (62.2%) | 0.7380 (57.4%) | ~60% |
| Swing time (s) | 0.4861 (37.8%) | 0.5477 (42.6%) | ~40% |
| Step length (m) | 0.5795 | 0.7080 | — |
| Stride length (m) | 1.2875 | — | 1.40 |
| Stride time CV (%) | 0.1185 | 0.0902 | <3% |
| Step length SI (%) |  | 19.9592% | 0% |
| Stride time SI (%) |  | 0.0311% | 0% |
| Double support (s) |  | 0.1598 | — |
| DSR (%) |  | 12.4252% | 20% |
| Cadence (steps/min) |  | 93.34 | 110 |
| τ_human hip (Nm) | 0.4375 | 0.4200 | — |
| τ_human knee (Nm) | 0.3825 | 0.4050 | — |
| PTF hip (%) | 34.31 | 36.52 | →100% |
| PTF knee (%) | 38.25 | 43.78 | →100% |
| S — Symmetry score | 0.0283 | → 1.0 |
| V — Stability score | 0.9752 | → 1.0 |
| B — Support score | 1.0000 | → 1.0 |
| E — Efficiency score | 0.7746 | → 1.0 |
| STR — Joint Strength | 0.3822 | → 1.0 |
| GRI — Gait Recovery Index | 0.5847 (58.5%) | → 1.0 |

# AI-generated Clinical Recommendations

CEMRR · Nazarbayev University · Robot-Assisted Gait Rehabilitation